Before Core Logging – A few key points

General

- Please submit a sample request and do so ahead of time!

- Make sure that the cores equilibrated to room temperature (recommendation: 5 h for IODP split liner; 10 h for GeoB split liner) – some parameters are T sensitive!

- Make sure that the core is positioned horizontally on the track.

- Split core halves: prepare the surface as flat as possible, remove mold. However, please keep in mind that you are dealing with archive halves in most cases!

- Be aware of that endcaps influence your measurements and adapt your procedure (see poster at the MSCL).

- Make sure to pick up the processed core sections from the end of the track (they will fall down otherwise)!

Linescanning

- Calibrate before you start measuring cores.

- To avoid reflections ensure that the surface is as dry as possible (careful preparation, e.g. with tissues).

- Support plate is available to allow continuous advance of the core sections on the track.

- Check quality of each image immediately after acquisition; it might be needed to calibrate twice a day.

GRA density

- Check if the lid at the Cs source container is open and latch it at the designated position.

- Calibrate once a day (first thing before you start).

- Check processed data immediately after acquisition (range within expected values?).

P-wave

- Make sure that there is good contact between transducers and core, where needed use water to improve contact (but do not flood the system).

- Check the sliding of both transducers regularly, if they got stuck: contact MARUM operator.

- Check the core flow, especially at the section breaks.

- Please always keep an eye on the P-wave amplitude.

- Check processed P-wave velocity data immediately after acquisition (range within expected values?).

Magnetic susceptibility

- Please Keep all metal pieces away from the sensor (watches, jewellery, belts, shoes, cans), these might disturb your measurement.

- Make sure that the point sensor is actually reaching the core surface when measuring.

Color reflectance

- Calibrate the spectrophotometer at the beginning (before running any core).

- Use suitable foil to cover the core.

Make sure that the surface is dry and clean.